Xitron Technologies, Inc.

BIS4000 Bio-Impedance Spectrum Analyzer System Utility

Version 1.00a

User Instructions

January 28, 1993

1. Minimum Hardware Requirements

PC compatible with at least 286 processor  ( 386/20MHz preferred )

Math co-processor

At least  640KB of RAM

MS-DOS 3.0 or higher

A RS-232C serial communication port, and the interface cable (supplied by Xitron).

2. Installation

Before you do anything, make sure that a copy of the original BIS4000 program diskette is saved ( make a backup disk ).  It is recommended that BIS4000 be run from your hard disk since BIS4000 will need disk space to store spectra, modeling, and body composition data.

a) Diskette Contents

The following files should be on the original program disk. If a file is

missing, please contact Xitron Technologies for correction.

README.DOC	This document in ASCII text format

BIS4000.EXE	Program

4000DRV.EXE	4000's or 4000B's  RS232 interface program

RS232DIA.EXE	RS232 port diagnostic program

XITRON.MSG	Program message file

COEFF.DAT	Factory coefficients set

b) Making a backup diskette

Insert the original BIS4000 program disk into drive A, At the DOS prompt, type:

diskcopy   A:  A:

Keep your original disk in the safe place. You might need it in the future.

c) Install the program on your hard disk

On your hard disk, create a sub directory for BIS4000 program.

After making the BIS4000 directory, insert the BIS4000 program disk into the floppy drive  A, and at the DOS prompt type:

copy   A:\*.*    C:\BIS4000

If no error occurs, the program installation is completed successfully.

3.	Program Startup

The correct DOS command syntax for BIS4000 program is BIS4000. To start up the program, simply type BIS4000 at the DOS prompt.

If no error occurs, the program should bring up the following disclaimer message:

The predictions of ECW, TBW, FFM, %Fat, and Fat Weight provided by this program are derived by equations and resistivity coefficients which are experimental and have not been thoroughly tested on a large number of subjects, or evaluated at the extremes of health, age or disease. Therefore, the results produced by this program should be used only for research purposes, and when possible, accompanied by appropriate criterion measures. Contact Xitron Technologies for information concerning the development and use of these equations.

Press any key to continue...

The program Main menu will appear if a key is pressed. If the program is attempted to be run and XITRON.CFG file does not exist, the program will ask the operator to first perform a setup. The program will hang up if the operator make a measurement without first performing setup. Therefore, it is recommended that setup be done prior to any other operation.

4.	Program Description and Usage

This program is menu driven providing the simplest user interface, and self-explained menu items. This section will describe the BIS4000 program operations, and how to operate the menus in the program. A menu has its title on the top, a list of menu items; the first item is always highlighted (selected) when a menu is first opened; arrow keys are used to change the menu item selection; return key is used to invoke the selected item. Each item has also a highlighted letter (cyan color) in its descriptive text; the highlighted letters in this manual will be shown in BOLD face letters; an item can be also invoked by entering its highlighted letter at the keyboard. ESC key can be used to exit a menu. The program will always check for: - disk drive readiness, bad file name, and file existence whenever it receives a new file name. If the drive is not ready, it will put up an error message and ask the operator for further corrective actions; if a bad file name was entered, it will assert a beep on the system's speaker and request for the new file name; if the file already exists, it will ask the operator to overwrite, append to the file, or cancel the operation.

4.1	Main Menu

After a key was entered at the above disclaimer message and initial setup is completed, the Main Menu should appear.

Main Menu

Communication Setup

Data Collection

Data Modeling

Volume Estimation

Complete Computation

Printing

Data File Manager

Quit (to DOS )

4.1.1	Communication Setup

This item is invoked if S is entered at the keyboard, or return is entered when this item is highlighted (from the Main Menu), or if a setup was never performed. When this item is selected, the following screen should appear.

Communication Setup

Xitron Impedance Analyzer Model	4000B

Com ( host ) Port	COM1

Baud Rate	19200

Printer Port	LPT1

Data Delimiter	TAB

Accept	Cancel

The following parameters, - Xitron Impedance Spectrum Analyzer model number, serial port, baud rate, printer port, are hardware dependent; therefore must be set up before any other operation. When the BIS4000 program first runs, it will look for XITRON.CFG file in the BIS4000 program directory; if it does not exist ( setup was never completed ), the program will ask the operator to perform a setup. When setup is selected, the BIS4000 will pull-up a menu of these hardware dependent parameters. Arrow keys are used to change the parameter selection; return key is used to initiate the parameter modification; ESC is used to cancel the modification to the selected parameter; highlighted letters  are used also to initiate the parameter modification.

a)	Analyzer Model

This is to tell the program which model of 4000 the program interfaces to. There are two different models:  4000, 4000B.

b)	Serial Port

This is to tell the program which serial port is used to interface with the 4000, or 4000B. If the operator does not know which serial port is connected to the instrument, the provided RS232 diagnostic program can help. Refer to Appendix C for instructions on this diagnostic program.

c)	Baud Rate

This is to tell the program which baud rate, the instrument's serial port is set at. The available baud rates for 4000 are: - 1200, 2400, and for 4000B are: - 1200, 2400, 4800, 9600, 19200. Default setting for 4000 is 2400 baud, for 4000B is 19200 baud. Refer to section "Front Panel Operation" chapter in the 4000 Operating Manual for instrument's serial port baud rate setting.

d)	Printer Port

This is to tell the program which printer port is used for printing. If the PC is not networked, LPT1 is always a default printer port. If the PC is networked, contact the network administrator for available printing resources.

e)	Data Delimiter

Since the output data might be imported into the spreadsheet for further study, it is required to use the correct data delimiter for different spreadsheet program. This is to tell the program which delimiter is used to separate the data fields in the output file. The available delimiters are: - Space, TAB, comma. Default setting is TAB.

4.1.2	Data Collection

This item is invoked if D is entered at the keyboard, or return is entered when this item is highlighted (from the Main Menu). This operation allows the computer control of the Xitron 4000, or 4000B Bio-Impedance Analyzer. This control extends the flexibility and ease with which data may be taken and processed. Reduced errors, ensured repeatability and extended programming are some benefits of using this module. Refer to "Accuracy Consideration" Chapter in 4000B Operating manual for further data collection information.

4.1.2.1	Input Parameters

The inputs to this module are: - data output filename, case number and the impedance data coming from the 4000, or 4000B through the RS232 interface, data delimiter.

- Output file name must be the correct DOS file format. If path name is included, the output file is stored under included path; if path name is not included, the output file is stored under the current directory.

- Case number can be any ASCII text string including spaces, and its length can be up to 80 characters.

- RS232 interface must be connected, and COM port, baud rate must be set prior to running this module. To diagnose the RS232 communication link, refer to Appendix C.

- Data delimiter is one of the 5 communication parameters defined above. This is used to separate the data fields within the data output file providing the spreadsheet importability. The default delimiter is TAB character which is compatible with Lotus 123, and Microsoft Excel. Refer to section 4.1.1 - Communication Setup for available delimiters and how to set it up.

4.1.2.2	Output File

The data collected from the 4000, or 4000B is saved to the data file upon the operator request. Each measurement is coupled with the operator specified case name, measurement date and time. A data file can keep more than one measurement. If the operator intends to use BIS4000 modeling module to model the data, the number of measurements per data file must be kept below 257. There will be no error generated if the number of measurements exceeds 256; however, the modeling module can load in only 256 cases.

For each case in the output data file, there are:

Case name	( can be any text of maximum 80 characters)

Date	( xx/xx/xxxx )

Time	( xx:xx:xx )

Frequency	Resistance	Reactance	Impedance	Phase

.	.	.	.	.

.	.	.	.	.

.	.	.	.	.

Next Case name

Date

Time

Frequency	Resistance	Reactance	Impedance	Phase

.	.	.	.	.

.	.	.	.	.

4.1.2.3	Data Collection Menu

The program will bring the Data Collection menu as shown below up when Data Collection item in the Main Menu is invoked.

Data Collection

Single Frequency

Single Frequency (continuous)

Dual Frequencies

Dual Frequencies (continuous)

Selected Spectrum

Selected Spectrum (continuous)

Log Spectrum

Log Spectrum (continuous)

Data File Manager

Quit (back to Main Menu)

a) Storing collected data

When one of the items in the Data Collection Menu is selected, the program will ask the operator if they wish to save the collected data to a file. If No is the choice, the program will prompt the operator to select the desired measurement frequency(ies). If Yes is the choice, the program will bring up the default output file name. The operator can modify this file name, enter the new file name, or press return to accept the default file name. After the output file name had been accepted, the program will display the default case number. The operator  can enter the new case number and press return, or simply press return to accept it. When the case number has been entered, the program will prompt the operator for the desired measurement frequency(ies).

b) Measurement frequency(ies)

The program will prompt the operator to input a frequency if the operator has selected single frequency option, two (2) frequencies if the operator has selected dual frequencies option, a set of selected frequencies or input frequency file if the operator has chosen the selected spectrum option, or the desired log spectrum if the operator has selected the log spectrum option. If at any time the operator does not wish to continue this process, press the ESC key.

c) Entering the selected spectrum

When Selected Spectrum measurement option is chosen, the program will give the operator the choice to load the frequencies from a file, or to input the selected frequencies from the keyboard. When the input and measurement are completed, the program will allow the operator to save these selected frequencies to file so that keyboard input is not required when the same set of frequencies are requested in the future.  There are some restrictions to the name of frequency file; it must only have extension ".FRQ" and each frequency file can only contain up to 512 frequencies. (logarithmic spacing is essential for modeling )

d) Selecting log spectrum

When Log Spectrum measurement option is selected, the program will bring up the following Log Spectrum menu:

Log Spectrum

25 Frequencies from 5KHz to 500KHz

25 Frequencies from 5KHz to 1MHz

50 Frequencies from 5KHz to 500KHz

50 Frequencies from 5KHz to 1MHz

Quit

Use arrow keys or highlighted letters to select the desired Log spectrum.

The four built-in Log Spectrums are:

1)-  	25 frequencies from 5K to 1M

5, 6, 8, 10, 12, 15, 19, 23, 29, 36, 45, 57, 71, 88, 110,

137, 171, 214, 266, 332, 412, 516, 644, 808, 1000

2)- 	 25 frequencies from 5K to 500K

5, 6, 7, 9, 11, 13, 16, 19, 23, 28, 34, 41, 50, 61, 73,

89, 108, 131, 158, 192, 232, 280, 342, 412, 500

3)-  	50 frequencies from 5K to 500K

5, 6, 7, 8, 9, 10, 11, 12, 13, 14, 16, 17, 19, 20, 23, 25, 27, 29, 32,

35, 38, 42, 46, 50, 55, 60, 65, 71, 78, 85, 93, 102, 111, 121, 133, 145,

158, 173, 189, 206, 226, 246, 270, 294, 322, 352, 384, 420, 460, 500

4)-  	50 frequencies from 5K to 1M

5, 6, 7, 9, 10, 12, 14, 15, 17, 18, 20, 22, 25, 27, 30, 33, 37, 41, 45, 50,

55, 61, 67, 74, 82, 91, 100, 111, 123, 135, 150, 165, 183, 202, 224, 248,

274, 302, 334, 368, 408, 448, 496, 548, 608, 672, 740, 820, 904, 1000

e) Performing measurement

If continuous mode is selected, the program will continue reading the measurements from the 4000 or 4000B, updating the data on the screen, and saving the measurements to output file if saving the data was requested. ESC is used to terminate the continuous measurement mode. If ESC is entered while 4000 or 4000B is in the process of performing a measurement, the program will continue to complete (and save the data to file  if requested) the last measurement before terminating. Warning: Continuous data collection can consume a great amount of disk space in a short period of time. Ensure that you have adequate data storage prior to invoking this option.

Prior to performing measurement it is recommended that the operator review the "Accuracy Consideration" chapter in the 4000B manual.

4.1.3	Data Modeling

This item is invoked if M is entered at the keyboard, or return is entered when this item is highlighted (from the Main Menu). This module is used to analyze data collected using the Xitron Technologies 4000, or 4000B Complex Bio-Impedance Spectrum Analyzer and the BIS4000 -Data Collection module. The software uses a proprietary iterative successive approximation technique to obtain the best possible fit to the standard cell suspension model which accounts for divergence in the material, and corrects for time delays in the conduction path. Refer to Appendix A for further details about modeling.

This software calculates the equivalent circuit component values in the model as explained in Appendix A. This is a preliminary fundamental step in either using Bio-Impedance Spectroscopy (BIS) to establish the relationship between these values and fluid volumes, or predicting fluid volumes after the relationship has been established.

4.1.3.1	Input Files

Impedance data file taken by Data Collection Module, or any data file having the format defined in section 4.1.2.2 - Output file.

4.1.3.2	Output Files

There are two output files having the same name as the input file and different extensions; these extensions are .MOD and .MDL. The *.MOD file contains the fit information such as:

- Case number	- ASCII text of 80 characters maximum,

- fit status,	- Correlation of fit using scalar Z, and

one of the following ratings:

0	Excellent fit

1	Good fit

2	Marginal fit

3	Questionable fit

4	Unable to obtain fit due to bad data

- Recf,	- in Ohm

- Ricf,	- in Ohm

- Cm,	- in nanoFarad.

- Td,	- in nanoSecs.

- alpha,

- Fc,	- in KHz

- DA	- in Degree

The *.MDL file contains the fit history and the comparative data between the measured impedance, phase and the predicted impedance, phase (based on the modeled parameters).

4.1.3.3	Operation

The program will bring the File Manager window up when Data Modeling item in the Main Menu is invoked. In this window, there are two different lists: - a list of filespec matching files, a list of available drives and sub directories. Use arrow keys to scroll the selections, TAB to switch the focus between these two lists. There are also other commands the operator can use by entering the following highlighted letters:

f	to enter the new filespec string,

a	to enter the new path string,

s	to get the file size, file date and time of the highlighted file,

v	to view the contents of the highlighted file,

e 	to pick the file for modeling,

q	to quit,

return	if the focus is in the file list, return is to select the highlighted file for modeling. If the focus is in the sub directory list, return is to change to the highlighted directory.

When a data file is selected, the program will launch another window. In this window, there is a list of all the cases contained in the selected data file. The first case number is always highlighted (selected) when this window is first brought up. As in the program menu, the operator can use the arrow keys to scroll through this list. Space bar is used to tag and untag a case number. Press A to start modeling all the cases. Press return to start modeling all tagged cases, or initiate modeling. ESC is used to go back to the main menu.

When modeling is initiated, the program will look at the list of tagged cases. If there is at least one tagged case, the program will sequentially fit all the tagged cases; If there is no tagged cases, the program will fit the current highlighted case.

While data fitting is in progress, the program will update the fit results on the screen in real-time. Data fitting can take considerable time, particularly on a 286 machine with corrupted data (typically a fit takes one minute on a 386 machine with good data). ESC can be used to terminate this procedure when neccessary.

If ESC is entered while fitting a list of cases, the program will stop fitting immediately and save the fitting data for all preceeding cases to the output files. Refer to section 4.1.3.2 for more details on the output files.

When the data fitting is completed, the program will ask the operator to view, save (or both) the fit status and fit data.

4.1.4	Volume Estimation

Please review the disclaimer on page 2 before proceeding. This item is invoked if V is entered at the keyboard, or return is entered when this item is highlighted (from the Main Menu). This module will estimate the ECW, TBW Volume, FFM and %Fat based on the following parameters: Gender, Recf, Ricf, Height, Weight and {SYMBOL 114 \f "GreekMathSymbols"}ecf, {SYMBOL 114 \f "GreekMathSymbols"}icf, Decf, Dicf. Refer to Appendix B for further information on these equations.

4.1.4.1	Input Files

There is one required input file - coeff.dat and the optional input spreadsheet file depends on which mode the operator is running. If single case mode is desired, the operator does not have to prepare the input information in the spreadsheet format since the program allows the operator to type in the input information in single case mode. However, multiple cases computation becomes useful when there are  many cases. Please refer to section 4.1.4.2.2 for information on multiple cases.

a)	Coefficient file	( coeff.dat )

This file contains the coefficients for the equations. The program will search for this file in the drive and directory which contains the BIS4000 program. For example: In our default case, if you install the program in  C:\BIS4000   directory, the coeff.dat file must be present in this directory as well. If the operator wants to use the new coefficients file, the new coefficients file must be copied over the factory supplied coeff.dat. In case you want to create your own  coeff.dat  file, you can use your favorite text editor and follow the format below:

{SYMBOL 114 \f "Symbol"}ecf-MALE, {SYMBOL 114 \f "Symbol"}ecf-FEMALE

{SYMBOL 114 \f "Symbol"}icf-MALE, {SYMBOL 114 \f "Symbol"}icf-FEMALE

Decf-MALE, Decf-FEMALE

Dicf-MALE, Dicf-FEMALE

Present Factory default  coeff.dat  :-

240, 206

824, 797

1.45, 1.48

1.31, 1.23

b)	Input Spreadsheet file

This file is needed if multiple cases computation is desired. The file name can be anything. It doesn't need to be in the same directory with the program.  If it resides in the directory other than the BIS4000's, the operator must specify the whole path name.  For example: If the spreadsheet file is on the disk in the floppy drive B, and its name is testdat.txt, you must enter B:\TESTDAT.TXT  as your input filename. The input file may be produced by the user using a text editor or from a spreadsheet program using a suitable ASCII text output format.  A maximum of  256  subjects are supported.  The data for each subject is expected in the following format:

Required fields

Subject ID	This is a "free format" ASCII text field, and is used to identify individual subject data; if there is any space or TAB characters, the double quotes are needed.

Gender 	This field should contain a numeric quantity indicating the gender of the subject. A zero (0) indicates male and a one (1) indicates female.

Height	This numeric filed contains the height of the subject in centimeters.

Weight	This numeric field contains the weight of the subject in Kilograms.

Recf	The Recf  result, in ohms, from modeling the measured spectrum.

Ricf	The Ricf  result, in ohms, from modeling the measured spectrum.

Optional fields

Vecf	Criteria extracellular fluid volume in Liters.  Will be ignored if set to zero or is not present.

Vtbw	Criteria total fluid volume in Liters.  Will be ignored if set to zero or is not present.

FFM	Criteria Fat-Free Mass in Kilograms.  Will be ignored if set to zero or is not present.

Fat Weight	Fat weight of the subject in Kilograms.

%Fat	Percent of fat within the total body weight.

Suggestion:  If the criteria data is not available, fill the

corresponding field with 0

4.1.4.1	Output Files

There are two possible output files - Output spreadsheet file, coeff.new (optional).

a)	 Output Spreadsheet file

Whose contents are the input descriptive, modeling results, optional criteria data, and five more fields added at the end of each record. They are predicted values of Vecf, Vtbw, FFM, Fat Weight, and Percent Fat.  If this module was invoked with  criteria data, these predicted values are calculated with the newly regressed coefficients. At the end of the output file, Three additional records are added with the subject ID of   "r",  "SEE",  "OFFSET" respectively. These records have all descriptive, modeling and criteria set to zero, and the prediction result fields contain the Pearson (r) - correlation, Standard Estimate of prediction Error, and the Mean Offset between the criteria and predicted data where computation is possible.

b)	New Coefficient file		( coeff.new )

This file is produced only if Multiple Case (with criterion) was invoked. It contains the new coefficients for equations after regressions.

You might want to update your  coeff.dat  file with the newly generated coeff.new file. It is recommended to keep the history of your  coeff.dat. Since coeff.dat history is important, the program automatically make a backup for it when the operator make changes or update its current contents.

Only 10 backups can be kept. The coeff.dat's  backups are file with extension from BK0 to BK9. If a coefficient file is updated, the program will rename the coefficient file to have the extension from BK0 to BK9 sequentially. The program derives the new extension by one increment from extension of the newest backup. The extension will roll back to BK0 if it reaches BK10. If more than 10 backups are needed, use different file name.

4.1.4.2	Operation

When Volume Estimation is invoked, the following menu will appear.

From this menu the operator can perform an estimation of ECW, TBW, FFM, Fat Weight and %Fat for single case, multiple case or multiple case with coefficient regression if criterion data is available. In order to invoke any item in this menu, the coeff.dat must be present in the same directory as the BIS4000 program.

Volume Estimation

Single Case

Multiple Case

Multiple Case ( with regression )

Coefficient Default

Quit (back to Main Menu)

4.1.4.2.1	Single Case

This item is invoked by entering S, or return at the keyboard when it is highlighted. This item is for single case computation with these  information available:   gender, height, weight, Recf, Ricf. When this item is selected, the following screen will appear.

{EMBED CDraw \s  \* mergeformat|}

Arrow keys and highlighted letters are used to select the data field for modification; ESC is used to discard the modification. The save button is for storing data to file; when this button is selected, the program will bring up the default output file name. The operator can modify this file name, enter the new file name, or press return to accept the default file name. Of course, any new file name received by the program will be checked for bad filename and file existence.

4.1.4.2.2	Multiple Case

This item is invoked by entering M, or return at the keyboard when this option is highlighted. This item is for multiple case computation with or without criterion data. If criterion data are present, the program will compute the estimations and the statistics between the criterion and predicted data. If the resistivity coefficients are not computed for the sample through the regression option, the predictions and statistics will be calculated based on the existing coefficients in the coeff.dat file. This option requires input data from a spreadsheet; so the operator must prepare a spreadsheet before running the program with this option. When this option is selected (from Volume Estimation Menu), the following screen will appear:

Input File:	Alt F

Output File:

There are two input fields: input spreadsheet file name, output spreadsheet file name.  Highlighted letters are used to select the input field; Alt F is used to pull up the  file manager. Refer to section 4.1.6 for information about the file manager.

When input file name is entered, the program will check for bad file name, and file existence. If there is error in the input file name, the program will keep asking the operator for the new file name. If no error occurs, the default output file name will appear. The operator can simply press enter to accept the default output file name, or modify this file name, then press enter. As soon as enter is received, the program will prompt the final message for confirmation. If the operator needs to make some changes, the highlighted letters are used to select an input field for modification.

The operator can also use file manager to pick the input file name. Alt F is  used to pull up the file manager. If the operator already in the editing mode, enter ESC before Alt F be typed in.

After both input and output file names have been selected, the operator can simply press return as the response to the confirmation message. The program then reads the data from the input file, calculates Vecf, Vtbw, FFM, Fat Weight, %Fat and stores the results in the output file. If criterion data is present, the program will compute the statistics (Pearson -r, OFFSET, SEE) between the criterion data and the predicted data and store the results in the output file as the last three records (Refer to section 4.1.4.1 for more details on the statistics used).Any error or warning will be shown on the screen as the operation progresses.

4.1.4.2.3	Multiple Case With Coefficient Regression

This item is invoked by entering C, or return at the keyboard when this option is highlighted. This operation is used for multiple cases and calculation of sample specific coefficients derived from criterion data. The predicted variables and the statistics(Pearson -r, OFFSET, SEE) will be calculated based on these new coefficients and stored in the output file as the last three records. If the criterion data is not present, the program  will only compute the predicted variables based on the resistivity coefficents in the existing coeff.dat file. The new coefficients are produced and stored in the coeff.new file only if the criterion data are available in the input file. The calculated body composition variables are stored in the output file.

Operating instructions for this option is identical to Multiple Case. Refer to section 4.1.4.2.2.

4.1.4.2.4	Coefficient Default

This item is invoked by entering D, or return at the keyboard when it is highlighted. This section is for coeff.dat management. When this item is selected, the following screen should appear.

coeff.dat

{SYMBOL 114 \f "GreekMathSymbols"}ECF Male   xxx.xx	{SYMBOL 114 \f "GreekMathSymbols"}ECF Female   xxx.xx

{SYMBOL 114 \f "GreekMathSymbols"}ICF Male   xxx.xx	{SYMBOL 114 \f "GreekMathSymbols"}ICF Female   xxx.xx

dECF Male   xxx.xx	dECF Female   xxx.xx

dICF Male   xxx.xx	dICF Female   xxx.xx

View	Replace	Save	Quit

XXX.XX represents the value for each coefficient from the coeff.dat file. If the operator wants to make any change to the coefficients, use the arrow keys or highlighted letters. In the section at the bottom, there are four buttons - v to View, r to Replace, s to Save, q to Quit.

View is for viewing the contents, and file data, time of any existing coefficient file. When V is pressed, a list of coefficient files will appear. The operator can scroll through this list by using the up, down arrow keys. To view a coefficient file in the list, highlight the file and press return. Once the return key is pressed, the contents of the highlighted file will be displayed. Use ESC key to get back to the previous screen.

Replace is for replacing COEFF.DAT with any modified coefficient file.  When this key is pressed, a list of coefficient files will appear. The operator can highlight the file to replace the COEFF.DAT, and press return. The program will ask for the confirmation. If the operation is confirmed, the program will make a backup of the COEFF.DAT file ( the back up file name is COEFF.BKX where X ranges from 0 to 9 ), then copy the highlighted file to the COEFF.DAT file. The COEFF.BKX file is for recovery purposes; In case the operator decides to go back to the previous coefficient set, this allows the operator to do so.

Save is for saving the latest changes to the COEFF.DAT file. When this operation is confirmed, the program will make a back up of the COEFF.DAT, then writes the coefficients up on the screen to the file COEFF.DAT.

Quit will bring the operator back to the previous menu.

4.1.5	Complete Computation

This item is invoked if C is entered at the keyboard, or return is entered when this item is highlighted (from the Main Menu). This operation is useful when the operator wants to measure, model, and  calculate the predicted values in one single operation. When this item is selected, the following screen will appear.

Complete computation ( single case )

Case	0001

Gender	Male

Height	0.00	centimeters

Weight	0.00	Kilograms

--------------------------------------

Recf	0.00

Ricf	0.00

Predicted Vecf	0.00

Predicted Vtbw	0.00

Predicted FFM	0.00

Fat Weight	0.00

Percent Fat	0.00  %

Start Analyzing	Save	Quit	Show (modeling data )

Use arrow keys or highlighted letters to enter or modify a data field. There are 4 other controls:  A to Start Analyzing, S to Save, Q to Quit, o to Show.

Start Analyzing is to initiate the measurement, modeling, and prediction process. When A is entered, the program will ask for selection of the type of desired log spectrum. Once the Log spectrum has been selected, the program will begin computation process. Warning: This complete computation option can consume considerable time if this program is running on the slow machine and poor quality data has been collected. Use the ESC key to abort this procedure if neccessary.

Show is to show the modeling data on the screen after Start Analyzing had been completed.

Save is to save the data to a file. The format of output file is described in section 4.1.4.1 - Output Spreadsheet file.

Quit is to go back to the previous menu.

4.1.6	Printing

This item is invoked if P is entered at the keyboard, or return is entered when this item is highlighted (from the Main Menu). Print manager and file manager are handled by the same module; therefore, they are the same except that some control buttons are different. File manager or print manager allows the operator to view or select a file from any directory. When the print or file manager is selected, a window will appear with two different lists: - a list of filespec matching files, a list of existing drives and sub directories. Use arrow keys to scroll the selections, TAB to switch the focus between these two lists. There are also other control buttons the operator can use by entering the following highlighted letters ( there are some control buttons not active in both file manager, and print manager ):

f	to enter the new filespec string,

a	to enter the new path string,

s	to get the file size, file date and time of the highlighted file,

v	to view the contents of the highlighted file,

e	to select the file,

d	to delete the highlighted file (not active in print manager),

c	to copy the highlighted file to other file (not active in print manager),

r	rename the highlighted file (not active in print manager),

p	to print the contents of the highlighted file to printer port,

q	to quit,

return	if the focus is in the file list, return is to select the highlighted file for modeling. If the focus is in the sub directory list, return is to change to the highlighted directory.

4.1.7	Data File Manager

This item is invoked if F is entered at the keyboard, or return is entered when this item is highlighted (from the Main Menu). Refer to section 4.1.6 for more detail.

APPENDIX A

Modeling

APPENDIX B

Volume Equations

APPENDIX C

RS232 Diagnostic

{PAGE|20}